## Slide 1
Xiaoyan Tian
Zhexuan Zhang
Stem Cell and Orthopedic Bioengineering Lab
Department of Biomedical Engineering
University of Miami College of Engineering
Mentor:
Dr. Chun-Yuh Charles Huang
BME 403 Report
December 2014
INSERT TITLE HERE
Phase Measurement System

## Slide 2: Objectives
To develop a system function in measuring the concentration of substance based on Ruthenium-oxygen quenching.
Notes: The oxygen-related substance can be oxygen, ATP, glucose, and some other substances that can alter the oxygen concentration due to different fabrication method of biosensors. So our project has a wide range of application.

## Slide 3: Objectives
Observation: biology-related labs
Problem: inefficient conventional concentration measurements; expensive devices.
Need: Fast, accurate, efficient, even portable concentration measurement
Notes: In biology-related labs, we may need to monitor some critical concentration change over time. In conventional concentration determination method, we have to draw some solution out, and then use titration, or some other methods depending on what we’re measuring, to determine its concentration, which takes too much time, and it can never show the concentration change over time, in other words, we need a system to monitor the concentration in real time ideally, or at least semi simultaneously.

## Slide 4: Review of Background
Fluorescence quenching of excited Ruthenium complex
Stern-Volmer Equation
Notes: The basic principle is based on the fluorescence quenching of excited ruthenium complex. Simply speaking, Ruthenium can be excited by blue light, and when it counters oxygen molecular, it would quench and emit a red light, which has a measurable decay time. And it was determined by the stern-volmer equation. In this equation, tao0 is a constant, Ksv is constant, so the oxygen concentration and the inverse of decay time has a linear relationship.
So the equation is only viable to oxygen concentration, but we can extend its viability by coating enzyme layer on top of oxygen sensing layer. For example, if we want to measure the glucose concentration, we can coat glucose oxidizer on top of ruthenium layer, the oxidizer will alter the concentration of oxygen which is a function of glucose concentration, then the oxygen are deliver to the ruthenium layer, and according to the decay time, the concentration of the oxygen can be determined, and the concentration of glucose can therefore be determined.

## Slide 5: Review of Background
Similar Product
NeoFox from Ocean Optics
Expensive
Required a PC connection
Notes: One of a similar product that uses this concept is the Neofox from Ocean Optics. But the whole system is about 3000$, and it need a PC to run its software in order to get its data out.
So what we’re going to do is to built a real portable and cheap device that has almost the same function.

## Slide 6: Design specifications
Must:
  The device must be able to acquire decay time from O2-Ru Quenching
  Oxygen concentration determination algorithm must be effective
Should:
    Quenching decay-time should be displayed on the LCD screen
  Device should be portable
Could:
  Calibration system could be incorporated into the microprocessor.
  The whole device could be integrated on a single chip.
Would:
  Touch screen would be more user-friendly.
  Higher resolution would be achieve by a more expensive microprocessor.
  A corresponding software would be developed.

## Slide 7: Material and Budget
Arduino Leonardo (Atmega 32u4) by Arduino Amazon.com $19.99
10 bit ADC with maximum sampling frequency of 77KHz.
Blue LED light by Bluecell Amazon.com $5.67/50PCs
Resistors, capacitors, wires, breadboard Amazon.com $17.62
Liquid Crystal Display Module Amazon.com $47.94
BPW34 Photodiode Arrow.com $4.84/10PCs
OPA2604 Operation amplifier digi-key.com $27.60/5PCs
Overall budget: $123.66

## Slide 8: Design Constraints
The product is much affordable than current devices.
Avalanche photodiode requires a high voltage, an over voltage protection will be incorporated in the device.
The whole system is based on a microprocessor, and it can be specifically designed for mass production
IEC 62304:2006 Medical device software
ISO 15223-1:2012 Medical device labeling
510(k)
Notes: 510(k): since similar device has already been in the market.

## Slide 9: Design Development
Functions that should be implemented:
Pulse generation
Signal acquisition
Decay calculation
Data display
Notes: Three out of four main functions of our project (pulse generation, decay calculation, and data display) relays on the microprocessor. At first, we chose arduino leonardo as the foundation of this project, since it’s language is based on C, and its own library is easy to learn, we thought it’s a good start. Next step is to ask the arduino to fullfill these functions.

## Slide 10: Design Development
Exponentially decay signal
Notes: Fist consider what is an exponential decay. It is after removing the excitation, the signal decays in an exponential fashion. So how do we generate such kind of decay. It can be simply done by using a square wave to excite the probe, and in every time point when the pulse goes off, the probe will return this kind of decay signal, ideally.

## Slide 11: Design Development
Matlab simulation 1 w/ 5us decay
Pulse generated by Microprocessor is 33Hz (30ms/cycle)
The range of decay time is from 2us - 5us
x-axis: 50ms/div
Notes: 33Hz is basically a random number, as long as the wave is square like and the time interval b/w each pulse is larger then 50us , the decay will appear. But in the figure we generated from matlab, it seems like that we can see no decay from the orange light signal. But consider that the time level in the graph is millisecond lvl, let’s zoom in to see what it really

## Slide 12: Design Development
Matlab simulation 2 w/ 5us decay
x-axis: 10us/div
Zoom in from last figure to see signal in micro-second level
Notes: When we zoom in to micro sec lvl, we can see the simulated decay clearly, and if we could acquire all data point on this decay curve, the decay time should be easy to calculate. So the next step is to see if our Arduino can record all these data points.

## Slide 13: Design Development
1st limitation: Arduino ADC sampling rate
http://arduino.cc/en/Reference/AnalogRead
100us  1 AD conversion
Notes: We went to search arduino’s database to find out it’s sampling rate. And very disappointing, we found that by using it’s own library, the sampling frequency is 10KHz, by that it means it will take 100 microseconds to complete an analog to digital conversion. Let’s go back and see how long it will take to bring the decay signal back to 0, it takes less than 10 microsends to decay to zero. This is a huge limitation arduino has, but we do not give up easily. We were thinking what is the arduino doing during this 100 microsecond period. Arduino does not tell us how it work, so we directly went to the heart of arduino, the real microprocessor on the arduino board.

## Slide 14: Design Development
Countering limitation 1
10KHz  77.8KHz by overriding Arduino library
fully understanding ADC principle
Notes: Reading the datasheet of the microprocessor is tough, but we manage to understand the ADC principle. For the sake of reducing power consumption, the arduino limits the ADC rate to 10KHz, but we can manuelly override the arduino setting, to accelerate the ADC sampling rate to 77.8KHz, that’s 13 microsecond per conversion. But still, after 13 us the signal has already diminished to zero. It looks like that it can’t be done with this microprocessor, but it’s not that simple. Let’s see this figure, it’s the ADC principle of this microprocessor. Basically, it has 3 time stage for an AD conversion. First stage is calling up the ADC function, telling that hey, be ready for taking the signal and convert it, 2 stage is piking up the signal and store it. Third stage is to digitize the analog signal. And we found out that the 13us process only convert the signal at exactly 1.5us after calling up this function. If we call ADC function at the moment when the signal began to decay, we will be able to catch the signal at 1.5us and it’s the most precious signal that can determine the decay time.

## Slide 15: Design Development
Countering limitation 1
10KHz  77.8KHz by overriding Arduino library
fully understanding ADC principle
Notes: We need two point to determine the decay time of an exponentially decaying signal. We’ve already got one as I discribed before. Another one is easy, since the value at the start of the decay is the same as the top stable signal, thus we can measure the stable signal in advance, and use that value as the value at time zero. Since stable value is relatively long 15 millisecond, so we actually took 100 stable value and average them to minimize the noise.

## Slide 16: Design Development
Arduino Language:
analogRead()
Code that write into the chip:
int adc_decay_start() //setup ADC7(A0) with normal conversion
{
ADCSRA |= (1<<ADPS2); //prescalar=16; ADC clk freq=1MHz;
//1 ADC clock = 1us;
ADCSRB |= (1<<ADHSM); //enable ADC high speed mode
ADMUX |= 0x47; //AVcc with external capacitor on AREF pin. right adjustment to maximize 10 bit.
//set ADC7 as input channel
ADCSRA |= (1<<ADEN); //ADC enable
ADCSRA |= (1<<ADSC); //ADC start conversion
//normal conversion does not manipulate ADCSRB
}
double adc_normal()
{
int sum = 0; //use to average normal_ADC reading
adc_decay_start();
for (int i=0; i<9; i++)
{
while(ADCSRA & (1<<ADSC)); //wait till ADC complete
sum = sum+ADC;
ADCSRA |= (1<<ADSC);
}
while(ADCSRA & (1<<ADSC));
sum = sum+ADC;
//ADCSRA &= ~(1<<ADEN); //don't need to shut down ADC
double normal_ADC_ave = sum/10.0/1023.0*1000*Vcc;
return normal_ADC_ave;
}
Notes: This slide shows a comparison b/w arduino language, and the code we wrote to directly access the microprocessor. Totally two different story. Arduino language is simple and intuitive, but it’s function is limited, by directly accessing the chip, we can do more things, but the language is not easy to understand.

## Slide 17: Design Development
2nd limitation: Arduino Pulse Generation
Task: call ADC function exactly at the moment when the excitation shuts.
Notes: In order to get the signal on the curve, we need to call ADC function exactly at the moment when the excitation shuts off. We need a really accurate Timer to set the pulse at exactly 30ms per cycle, which is 30,000micro second per cycle. But the arduino’s timer resolution is 4mciroseconds, which means that calling up it’s timer function, it takes 4 microseconds. This 4 microsecond could interfere with calling up the ADC function. So this is the 2nd limitation that we have to counter.

## Slide 18: Design Development
Arduino language:
while (1)
{
unsigned long currentMicros = micros();
if(currentMicros - previousMicros > interval)
{
previousMicros = currentMicros;
if (ledState == LOW)
ledState = HIGH;
else
ledState = LOW;
digitalWrite(ledPin, ledState);
}
}
Code that write into the chip:
TCNT1 =0;
while (LED_flag ==0) //LED dark
{
if (TCNT1 >= 33332) //counting
{
PORTB |= (1<<5); //LED light up
TCNT1 = 0; //Reset timer value
LED_flag = 1;
}
}
while (LED_flag == 1) //LED light up
{
if (TCNT1 >=20000 & TCNT1 < 20020)
normal_ave = adc_normal();
else if (TCNT1 >= 33332)
{
PORTB &= ~(1<<5); //LED shut down
LED_flag =0;
TCNT1 = 0; //reset timer value
}
}
Notes: Again, the arduino language is not so sufficient, so we turn to the datasheet of the chip. We found that the chip has a counter that has a resolution of 1us, and we don’t have to call this function to read the time, it’s always counting the time. In other words, calling the chip’s Timer is instantaneous, not like arduino, need 4microsecond to read the time. This is the comparison of wrting code to fulfill similar function.

## Slide 19: Design Development
Data Display Arduino built-in LCD library
lcd.print();
Notes: Finally, the last part of the function that auduino should achieve is Data display. And this is the only part that we took the advantage of arduino, which is simple and intuitive. We just connect the circuit that provided on it’s website and one line of code, the data can be display on the LCD screen.

## Slide 20: Design Development
Notes: The basic block diagram of our project is this.
So we went through the main functions of our microprocessor, and it’s time for us to see if it really works. Last semester, we used RC circuit to simulate the exponentially decaying signal, and test if the signal can be pick up by the microprocessor and the decay time can be calculated by the microprocessor.

## Slide 21: Proof of Concept
Notes: We just simply use RC circuit as out test subject to see if the microprocessor works.

## Slide 22: Proof of Concept
Student t-test b/w group 1&2
t=-65.7
Std=0.0122
The probability of they’re paired is 0.000
Student t-test b/w group 1&3
t=-122
Std=0.0137
The probability of they’re paired is 0.000
Student t-test b/w group 2&3
t=-84.9
Std=0.0102
The probability of they’re paired is 0.000
Notes: We used three different capacitor and record 10 data for each capacitor, and run a student t-test b/w each group. It turns out that each data set are independent, and each RC circuit has it’s own decay time.

## Slide 23: Proof of Concept
Prove the one-to-one relationship between decay-time and capacitance
Our microprocessor and algorithm is viable.
Coefficient difference may due to capacitor error

## Slide 24: Proof of Concept
Using Silica photodiode to test if algorism is viable
Mean: 1.5211
Std: 0.0137
Notes: The RC circuit directly gives us electric signal, but our ultimate goal is to test the light decaying signal, so we tried to use a simple photodiode to test if we can recover some light signal before we going further. We used a silica photodiode to examine if we can measure the decay of the blue light signal. The answer is yes. The same circuit connection, and program gives us a stable value for the blue light decay. But the problem using the silicca photodiode is that it cannot detect the orange light signal coming back from the probe.

## Slide 25: Design Development
Phase 2
Avalanche Photodiode
Hamamatsu APD module C12702
Short wave type
Response time: <1ns

## Slide 26: Design Development
Notes: Then, with the help of Dr. Montero, we migrate our circuit from bread board to a shield that makes our project more compact

## Slide 27: Design Development
False signal due to power consumption
LED drawing too much power from power supply, the ADP was not powering up properly
Notes: After we received the APD, we connect our whole product to the probe and the APD, we got this nice signal on oscilliscope. But unfortunately, we test the probe under high oxygen conc solution and low oxygen conc, we found out that the decay time showing on our LCD does not change. We was confused, and went back to see our code to see if there’s anything wrong in the code. By accident, we disconnect the probe and the APD, and we found out that the APD is still sending us back the same signal. Then we turn to Dr. Bohoquz, and he found out that it’s because the LED drawing too much power from the power source when it lights up, so accordingly, the APD is actually turned off when the led lights up, and turn on when the led shuts off. Dr. Montero provides us a stable power source and solve the problem, now LED and APD will work properly at the same timme.

## Slide 28: Design Development
3rd limitation: APD will only give us peaks instead of square wave when we’re using 33Hz light pulse
APD has a 4KHz HPF inside its module, which can eliminate other light from environment, but at the same time eliminate our 33Hz signal too.
Matlab simulation 3
Notes: Instead of giving us square wave, it gives us spikes, as the green line looks like.

## Slide 29: Design Development
3rd limitation: APD will only give us peaks instead of square wave when we’re using 33Hz light pulse
Notes: Let’s zoom in to see what it really looks like in micro-second lvl.

## Slide 30: Design Development
3rd limitation: APD will only give us peaks instead of square wave when we’re using 33Hz light pulse
Arduino cannot measure negative voltage.
The relationship b/w decay signal and APD signal is obscure.
5KHz
354KHz
512KHz
Notes: It’s very difficult to determine the decay time even we have the APD signal like that. And the distortion of the signal is mainly because of the high pass filter. We try to use high frequency blue light pulse instead of 33Hz to eliminate the HPF distortion. We connect LED to function generator, and APD to oscilloscope, and see at what frequency lvl, can the photodiode recover exactly the original light signal. We tried 5KHz, 300KHz, and 500KHz, and we can see from the pictures, when frequency is at 100KHz lvl, the signal can be recovered well. So Dr, Bohoquz suggest us a new way to see if we can get the decay time.
We went to Dr. bohoquz again, and he provide a new method to accommodate the APD’s high pass filter.

## Slide 31: Design Development
New Design
100KHz light signal
Matlab simulation 4
Notes: We use 100KHz, and use matlab to simulate the behavior of orange light and APD signal.

## Slide 32: Design Development
New Design
100KHz light signal
Matlab simulation 5
Notes: 5us decay time simulation

## Slide 33: Design Development
New Design
100KHz light signal
Matlab simulation 6
Notes: 5us decay time simulation
2us decay time simulation

## Slide 34: Design Development
New Design
100KHz light signal
Matlab simulation 7
Notes: 5us decay time simulation
2us decay time simulation

## Slide 35: Design Development
New Design
100KHz light signal
Matlab simulation 8
Notes: LUT

## Slide 36: Design Development
4th limitation: When generating such high frequency, Arduino may miss some waves.
Notes: Actually, the missing waves are quite a lot, and completely random, it means that it can severely distort the output signal we can get.

## Slide 37: Design Development
4th limitation: The APD still can not detect orange light signal even though we can see the light
Notes: Actually, the missing waves are quite a lot, and completely random, it means that it can severely distort the output signal we can get.

## Slide 38: Proof of Concept
Chart: Phase Measurement Device- Pareto Analysis
| Concept development | Sensor selection | Timer and ADC Coding | Circuit construction | System perfection | Testing | Decay-time Algorithm | LCD coding
Parameters | 30 | 20 | 10 | 10 | 10 | 10 | 5 | 5
Cumulative % | 0.3 | 0.5 | 0.6 | 0.7 | 0.8 | 0.9 | 0.95 | 1
Cut off % | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8
Notes: As for now, in the middle of the project, we have been testing our decay time algorithm. This is how we implement the algorithm concept. We simply test the decay time of the LED capacitance.

## Slide 39: Future
Theoretically, old and new design can both work if we have the most appropriate sensor and microprocessor.
Faster microprocessor frequency
More sensitive APD

## Slide 40: Acknowledge
We would like to thank Dr. Huang, Dr. Salas, Dr. Montero, Dr. Bohorquez for their help during our project.

## Slide 41
INSERT TITLE HERE
Thank you for your attention!